BioPlex Training Overview

Rev 05/06

Luminex xMAP technology allows for the simultaneous analysis of multiple analytes in samples- including cytokines, total proteins, and phosphoproteins. The principle of the assay is similar to a capture sandwich immunoassay- an antibody to the target protein is covalently coupled to internally dyed beads. When the beads are incubated with sample, the protein of interest is captured and then a biotinylated antibody for a different epitope is added to the reaction, which is then detected with streptavidin-phycoerythrin. Using a dual-laser flow-based reader, beads are analyzed for the detection antibody and the internal bead signature, identifying both the protein analyzed and the level bound to the bead.

The Lauffenburger lab has a BioPlex, which is the BioRad version of the Luminex platform.  The general policies for this are outlined in this document and must be followed.  This is an expensive piece of equipment with expensive reagents. Failure to follow the rules can result in loss of priveleges to use the instrument.

Who can use the BioPlex?

- Any graduate student/post-doc in the Lauffenburger/Griffith labs

- Lauffenburger/Griffith UROPs at the discretion of Doug and their individual mentors

- Outside users at the discretion of Doug and in cooperation with the Lauffenburger/Griffith users

Regardless of above status, you MUST be trained by the BioPlex Czar

Reagents for the BioPlex

Reagents consist of kits for specific analytes, buffers, cell lysis buffers, calibration kits, and filter plates.

- Individual kits are your responsibility to buy.  Discuss with the BioPlex Czar how to get our account specific pricing for BioRad kits.  Also, keep aware of overlap of your kit needs with others in the lab- there may be price breaks for buying bulk packs that you can split.

- Buffers are bought in bulk to get a price break and are for DAL/LGG lab members only.  These are ordered by the BioPlex Czar and are found in the cold room area near the Western blotting materials.

- Cell Lysis Buffer is bought in bulk to get a price break and are for DAL/LGG lab members only. This is ordered by the BioPlex Czar and components are found in the cold room area and door of -20oC freezer next to -80 oC freezer...

- Calibration Kits/Sheath Fluid are shared by all users and bought by BioPlex czar.  The calibration kit can be found in cold room.  You will not be allowed to change the sheath fluid, so don’t worry where it is!

- Filter Plates may be company specific.  For BioRad kits, there are plates in the buffer kits and additional plates are ordered from Millipore (cat # MSBVN1250) – these are for DAL/LGG lab members only.  They will be ordered by the BioPlex Czar and are found in the cabinet near the BioPlex..

For all common reagents – help everyone out by letting the BioPlex Czar know 	when things are running low or if you are going to need a lot of something in the 	future.  Don’t expect instant replacements on things you fail to report!

Instrument Startup/Use/Shutdown

General

- Sign up for use of instrument on Google calendar – estimated time is 1 hour for startup/shutdown procedures and 30 minutes per plate to read.

- In addition, fill in instrument log next to the BioPlex

- Each user has there own folder for protocols and result files- these are in MyDocuments\BioPlex Users\last name.  It is recommended that you back up your data frequently to another location!

- There is an area set up with the vacuum manifold and plate shaker to do assays to the left of the BioPlex reader.  Please keep this area clean, including emptying vacuum flask as needed and wiping out the vacuum manifold.  As this is by the window, be sure to cover your plate during all incubations to prevent photobleaching.

Start Up Routine – (40 minutes)

- Turn on instrument using the power strip switch.

- Turn on computer if not already on.

- Open BioPlex Manager 4.1

- On Quick Guide menu, click on ‘Start Up’

- Follow instructions using MCV plate.

- Laser takes 30 minutes to warm up before you can proceed.  To check time remaining select ‘Instrument’ on top tool bar, then ‘Instrument Info’ and look in Device Status.  Can also look at status bar at bottom of screen.

- Once laser is ready, on Quick Guide menu, click on ‘Calibrate’

- Follow instructions using Calibration Kit (in cold room) and MCV plate.

  - Enter your name

  - Match calibration values to the control number on each tube.  New kits will be added by the BioPlex Czar to the list.

  - Always do CAL 1 – this calibrates the bead classification.

  - CAL 2 – calibrates reporter fluorescence. For phosphoprotein assays select PMT HIGH; for cytokine generally select PMT LOW for CAL 2.

  - Be sure to vortex each calibration vial for 30s before use.

  - Calibration is OK for changes of +/- 2oC or until shut down.  The software will tell you if calibration needs to be redone.

  - If you run a plate at one PMT setting and your next is at the other PMT setting, be sure to redo the CAL 2 calibration.

Reading plates

- Set up your protocol by selecting ‘New Protocol’ in Quick Guide or ‘Open Protocol’ if previously used.

- Details can be found in the BioPlex Manager 4.1 manual, but here is a quick outline:

  - Describe/format plate:  indicate number & direction of replicates, mark wells as Blank, Standard, Unknown (X), or Controls

  - Copy and paste your unknown descriptions from a previously made column in Excel

  - Enter standards information if doing cytokines (kit should provide concentrations of the standard curve you made) or controls if doing phosphoprotein assays.

- Run Protocol

  - Select bead region map (usually 100)

  - Set number of beads (recommended is 25 for phospho assays, 100 for cytokine assays)

  - In Advanced Settings, check that DD Gates match the kit instructions and any other kit specific changes.  Normally want ‘pause if errors’ unselected, and ‘autosave’ selected.

  - Click Start- give your results file a name (and be sure it is placed in correct directory).

  - Data is also saved in the protocol file- allowing reanalysis if you don’t properly set up plate.  Therefore, if you use a previously set up protocol, do ‘Save As’ to prevent overwriting raw data file.

NOTE:  most parts of the protocol set up can be changed after the plate is run and reanalyzed.  You absolutely MUST correctly tell the number of wells to assay and pick the correct analytes however.

If reading multiple plates, or another user is planning to use in less than 2 hours…

- On Quick Guide menu, click on ‘Wash between plates’

- Follow instructions using MCV plate.

Shut Down

If done and no other user is signed up for next 2 hours, you must shut down the instrument.  THIS IS CRITICAL FOR KEEPING THE BIO-PLEX RUNNING.. most problems are a result of clogging of the fluidics resulting from a failure to flush and disinfect before idle periods.

- On Quick Guide menu, click on ‘Shut Down’

- Follow instructions using MCV plate- takes 10 minutes.

- Rinse MCV plate with MilliQ water and invert to dry on KimWipe.

- Shut down BioPlex Manager software

- Shut down instrument using power strip.

Troubleshooting

- In protocol run advanced settings you can opt to pause run for various errors.  If errors occur you will then be prompted to do “Remove Air Bubbles” or “Unclog” functions on the top bar.  Follow the instructions.  The assay can then be continued in “Rerun/Recovery” mode in the protocol.. you select the checkboxes of wells that need to be rerun.

- If a needle can not be unclogged by these functions, it needs to be changed out and thoroughly cleaned.  This must be done by either the BioPlex Czar or a user who is authorized for Troubleshooting.. the list of these users is in the binder by the BioPlex.

- Monthly validations of the optics and fluidics will be done by the BioPlex Czar.  If you notice a lot of problems especially with bead classification, please let them know and the validation can be done as a check.